## Slide 1: Application of innovative technologies in PET imaging
Dr Grzegorz Korcyl
Inż. Karol Farbaniec
Gabriel Hajduk
Department of Information Technologies
Jagiellonian University in Cracow
3rd Jagiellonian Symposium on Fundamental and Applied Subatomic Physics
The work is supported by the Foundation for Polish Science by grant no. TEAM/2017-4/39 and the Polish Ministry for Science and Higher Education grant no. 7150/E-338/M/2018

## Slide 2: Key concepts
Dynamic evolution of Field Programmable Gate Arrays
  FPGA, CPU, GPU in a single, compact package
  Novel development methodologies
  Applied in Machine Learning, Video processing, High Performance Computing etc.
Exploration of true real-time image
reconstruction mechanisms
High-level processing in real-time
  Shift processing algorithms from offline CPU/GPU to online acquisition system hardware
  Construct Event-By-Event processing pipeline
G. Korcyl, P. Korcyl, “Investigating the Dirac operator evaluation on FPGAs”,
Supercomputing Froniters and Innovations, 2019
Xilinx.com

## Slide 3: General scheme
Hierarchical architecture
  Endpoints
    Digitizers
    Buffers
  Concentrators
    Collect data from several Endpoints
  Controller
    Controls the readout
    Collects data from all Concentrators
Two communication directions
  Upstream
    Synchronization
    Control/monitoring requests
  Downstream
    Readout data stream
    Control/monitoring responses

## Slide 4: J-PET
192 scintillator strips
1536 analog signals
32 TDC Endpoints
8 Concentrators
1 Controller
  Processing pipeline on Zynq 45
    8 paralel data streams
    Application of calibrations
    Search for geometry and timing coincidences
    LOR and ROR calculation
    Visualization by Linux on integrated ARM core
  Resource utilization: 75%
  Performance: 42 Mhits/s
  Power consumption: 20 W
G. Korcyl, et al., “Evaluation of single-chip, real-time tomographic data processing on FPGA -SoC devices”,
IEEE Transactions on Medical Imaging, DOI: 10.1109/TMI.2018.2837741, 2018
B. Flak, „Data visualization system for PET device”, Master thesis, AGH Kraków, 2018

## Slide 5: Modular JPET Readout
 | JPET | Modular JPET | 
Scintillators | 192 | 312 | 1.6x
Analog channels | 1536 | 4992 | 2.9x
Digitizers | 32 | 48 | 6x
Logic [k cells] | 350 | 5400 | 15.4x
Memory [Mb] | 19 | 272 | 14.3x
DSP | 900 | 3972 | 4.4x
ARM cores | 2 | 4 + 2x RT + 1x GPU | >2x

## Slide 6: Continuous readout
Controller
Concentrator
Concentrator
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
Utilization: only 10 % of resources
Korcyl G. et. al.: A system for acquisition of tomographic measurement data
Patent Nr: US10007011B2, 2018, WO/2015/028594
48 endpoints 12 per concentrator
Synchronous Aurora links for internal com.
50 khz readout reate
10 Gb Ethernet raw data output

## Slide 7: Full processing pipeline
Controller
Concentrator
Concentrator
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
010101
Calibrations
Calibrations
Interaction points
Interaction points
Coincidences
LOR calculator
Sinogram
generation
Sinogram
filtration
Image
reconstruction
Numerically verified, simulated on NEMA IQ phantom data
Chart: Level of parallelization
| 1 | 2 | 4 | 8 | 16
resources | 28.7852 | 30.5855 | 34.1707 | 41.337 | 55.6634
refresh | 286.277 | 572.554 | 1145.1081 | 2290.2162 | 4580.4324
Sinogram processing pipeline (e.g. 128x128 @ 200 MHz)
Final image generation and visualization on ARM/GPU
K. Farbaniec, „Tomographic data processing and visualization on programmable devices”, UJ, Master thesis, in preparation
Raw data path preserved
2.8 Ghit/s

## Slide 8: Summary
We have demonstrated that FPGA-based hardware is suitable for high-level tomographic data processing
  First prototype can process data streams from the current JPET scanner
  Naive reconstruction and data visualization on a single device
  Processing pipeline flexible of unusual event selection on hardware level
Hardware processing and Modular JPET
  Scalable system, flexible geometry and event selection
  Expands the spectrum of applications
FPGA devices emerge in many sectors of science
  High computing potential, compact, power efficient
  Just like GPUs few years ago
  We are exploring their capabilities for PET
G. Hajduk, „Wizualizacja danych tomograficznych przy użyciu gogli holograficznych”, PK, Master thesis, in preparation