## Slide 1: Synthesis and characterization of the plastic scintillators for the total-body J-PET scanner
Łukasz Kapłon
Faculty of Physics, Astronomy and Applied Computer Science
Jagiellonian University
28.06.2019

## Slide 2: Outline
Chemistry of plastic scintillators
Optical properties measurement:
- light yield
- emission spectra
- decay time
Technical attenuation length of long plastic scintillator strips

## Slide 3: Motivation
Develop plastic scintillators with high light transparency for J-PET construction.
Measure properties of commercial scintillators and choose the best for prototypes building.

## Slide 4
Scintillation Principles: light emission mechanism

## Slide 5
Scintillation Principles: chemical components

## Slide 6
Scintillation Principles: cell casting polymerization
1. fast heating from room temperature 20 °C to 60 °C in 1 hour, 40 °C/h,
2. slow heating from 60 °C to 180 °C in 60 hours, 2 °C/h,
3. polymerization in 180 °C during 20 hours,
4. cooling from 180 °C to 90 °C in 30 hours, 3 °C/h,
5. first annealing in 90 °C for 4 hours,
6. cooling from 90 °C to 60 °C in 10 hours, 3 °C/h,
7. second annealing in 60 °C for 4 hours,
8. cooling from 60 °C to 20 °C room temperature in 10 hours, 4 °C/h.

## Slide 7
Scintillation Principles: light yield of scintillators
Amax
Amax scintillator – middle of Compton edge for sample
Amax stilbene – middle of Compton edge for sdandard
Light yield (LY) is a conversion efficiency of radioactive energy deposited in scintillator into light and is expressed in light photons
generated per megaelectronvolt of deposited energy (photons/MeV) by radiation quanta or particle. Comparing light yield of produced plastic
scintillators to light yield of anthracene or stilbene crystal scintillators is known as relative light yield.

## Slide 8
Scintillation Principles: emission and absorption spectra
Absorption (green) and emission (red) spectra of two fluorescent substances commonly used in plastic scintillators
Emission spectrum of plastic scintillator with p-Terphenyl and POPOP
p-Terphenyl emission spectrum overlaps with POPOP absorption spectrum
p-Terphenyl + POPOP
1-4 wt. %
~0.05 wt. %

## Slide 9
where following symbols denote:
K - normalization constant
Θ – time of the gamma quantum interaction with scintillator
t - time
td - decay time of the final light emission
tr - average time of the energy transfer to the wavelength shifter
σ - standard deviation reflects total uncertainty of light emission
from scintillator and TCSPC components: excitation diode,
photomultiplier tube and electronics.
Scintillation Principles: light pulse properties
Decay time can be defined as time over which the light signal from scintillator decreases to 1/e of its maximum value.
To measure decay time of plastic scintillators time-correlated single photon counting (TCSPC) was used.

## Slide 10: Results
Chemical composition | Light yield [photons/MeV] | Decay time [ns] | Emission maximum [nm]
Polystyrene 2% PPO 0.03 POPOP 0.03 bis-MSB | 11200 | 1.8 | 423
Polystyrene 2% BPBD 0.06 POPOP | 10770 | 1.51 | 422
BC-420 polyvinyltoluene | 10000 | 1.5 | 391
EJ-200 polyvinyltoluene | 10000 | 2.1 | 425

## Slide 11: Commercial plastic scintillators

## Slide 12: Technical attenuation length
The aim of this measurement was to determine technical light attenuation length of plastic scintillator strips used for J-PET modules construction.
The technical attenuation length (TAL) of a plastic scintillator bar is defined as the length of scintillator reducing the light signal by a factor of e and depending upon:
bulk transmission of the scintillator,
scintillator thickness,
shape of scintillator,
reflective property of the surfaces.
500 to 1000 mm long scintillators
+ y0
A1 – amplitude
x - distance from light detector
t1 - technical attenuation length
y0 - background

## Slide 13: Plastic scintillator strips used in J-PET
 | BC-420 | EJ-230 | BC-404 | EJ-200 | UPS-923A | SP32 | Epic PS
Dimensions [mm3] | 5x19x300 | 7x19x500 | 6x24x500 | 6x24x1000 | 6x24x1000 | 6x24x1000 | 6x24x1000
Used in J-PET prototype | 24 strips small barrel with PMTs | 192 strips big barrel with PMTs | 312 strips modular PET with SiPMs | 1m long with WLS and SiPMs | Research for one meter module with SiPMs |  | 
Data from manufacturers |  |  |  |  |  |  | 
Decay time [ns] | 1.5 | 1.5 | 1.8 | 2.1 | 3.3 | 2.5 | 2.4
Light output [% of anthracene] | 64 | 64 | 68 | 64 | 56 | 56 | 50-60
Wavelength of maximum emission [nm] | 391 | 391 | 408 | 425 | 418 | 425 | 415
Light attenuation length [cm] | 110 | 120 | 160 | 380 | 400 | NA | 200
Polymer base | PVT | PVT | PVT | PVT | PS | PS | PS

## Slide 14: Overlapping of 365 nm UV lamp emission spectra and scintillator absorption
UV light

## Slide 15: Setup for TAL measurements
In measurement following equipment was used:
UV lamp with 365 nm wavelength of maximum emission;
silicon PIN photodiode BPW 34 B with 2.73x2.73 mm2 active area;
digital multimeter Rigol DM 3064;
power supply ZUE 5352 M;
aluminium rail with 8 mm diameter holes drilled every 50 mm.

## Slide 16: Summary results

## Slide 17: Summary results
Scintillator 6x24x1000 mm^3 | TAL for LR scan [cm] | TAL for RL scan [cm] | TAL from specification [cm]
EJ-200 PS56 | 145.74 ± 11.65 | 149.59 ± 4.09 | 380
EJ-200 PS57 | 163.55 ± 7.84 | 145.65 ± 4.67 | 
UPS-923A PS1 | 62.59 ± 1.78 | 61.26 ± 0.85 | 400
UPS-923A PS2 | 69.45 ± 1.91 | 66.63 ± 0.76 | 
Epic PS1 | 32.25 ± 1.73 | 36.17 ± 0.93 | 200
Epic PS2 | 41.11 ± 1.56 | 41.25 ± 1.04 | 
SP32 PS1 | 22.67 ± 0.76 | 23.22 ± 0.60 | NA
SP32 PS2 | 24.79 ± 0.67 | 23.92 ± 0.64 | 

## Slide 18: Summary
Plastic scintillators with the different chemical compositions were synthesized.
Time-temperature cycles were established for polymerization in small cylinders as well as for polymerization in the glass mold allowing to manufacture plastic scintillator strips.
Spectroscopic and optical properties of these polystyrene scintillators were measured. Relative light output, decay time, emission spectra were measured to develop best composition.
Technical attenuation lengths were measured to choose the best scintillator for J-PET construction.

## Slide 19: Thank you