## Slide 1
Physics & Instrumentation Group
Department of Radiology, University of Pennsylvania
Performance characterization and human imaging experience with the long axial field-of-view PennPET Explorer Whole-body Imager
S Surti1, V Viswanath1, MJ Parma1, ME Werner1, JP Schmall1, MJ Geagan1, G Muehllehner2, JR Reddin1, AE Perkins3, ME Daube-Witherspoon1, JS Karp1
1University of Pennsylvania, Philadelphia, PA
2KAGE Medical, Philadelphia, PA
3Philips Healthcare, Cleveland, OH
3rd Jagiellonian Symposium: June 27, 2019
Funding support:
Department of Radiology, U Penn
Philips Healthcare
NIH R01-CA113941
NIH R01-CA206187 (Explorer consortium)
NIH R01-CA225874
NIH R33-CA225310
Notes: Discuss the design of a new scanner with large axial FOV
Acknowledge the many people that have contributed to this project

## Slide 2
Status of PET technology and potential for improvement
  Previous generation of PMT-based PET scanners achieved system timing resolution of > 400ps and > 4 mm spatial resolution
    Lu-based scintillators
    Light sharing detectors with large PMTs
  Advent of SiPMs has led to the development of digital PET
    Improved timing resolution (200-400ps)
    Some improvements in spatial resolution

## Slide 3
Status of PET technology and potential for improvement
Pre-therapy 616 s scan
Siemens Biograph mCT
Sp. Res.: 4.2 mm
Timing res.: 538 ps
Sensitivity: 10 kcps/MBq
Point source sens.: 6.7%
Post-therapy 426 s scan
Siemens Biograph Vision
Sp. Res.: 3.6 mm
Timing res.: 216 ps
Sensitivity: 15.6 kcps/MBq
Point source sens.: 8.3%
Notes: Same patient first scanned on mCT (12.4 mCi, uptake: 51 min, CBM, scan duration: 616 s) and 6 months post-treatment scanned on the Vision (15.0 mCi, uptake 66 min, CBM, scan duration: 426 s) showing disease progression.
Both used a 5 mm Gaussian post-reconstruction filter.
Despite a shorter imaging time, reduced noise and sharper rib definition in the Vision image visible due to improved physical performance.

## Slide 4
How to improve PET system sensitivity?
Thicker crystals or long AFOV?
Diminishing gains with crystals > 3cm thick
Point source sensitivity ∝ AFOV
Long AFOV also images more of extended objects (∝ AFOV)
Total effective sensitivity ∝ AFOV2
Sens. ∝ AFOV
Point source sensitivity
Surti, et. al., PMB 2013, 3995-4012

## Slide 5
How does object attenuation impact PET system sensitivity?
Point source in 20-cm diam. cylinder
27 cm
20 cm
10 cm
Axial field-of-view (cm)
0 20 40 60 80 100 120
Diam. =
Gain 2.3
Gain 2.5
70 cm
140 cm
AFOV
20 cm
Uniform activity in 20-cm diam.
Gain 9.0
Gain 25
Gain 40
200 cm
Scanner length
  Longer axial FOV increases sensitivity – increased solid angle in 3D
  Gain in point source (per slice) sensitivity limited due to attenuation
Notes: Brief review benefits of large axial FOV
Need to understand the difference of benefit gained per axial slice (point source gain) vs total body

## Slide 6
Advantages of long AFOV PET
  Increased sensitivity, enabling:
    Lower injected activity (or dose) – useful for pediatric imaging, serial imaging for treatment response, screening, etc.
    Shorter scan times – reduced patient motion and possibly breath-hold imaging
    Delayed imaging – better biological contrast
    Higher temporal resolution – improved estimation of TACs in dynamic imaging
    Imaging of low-yield or long-lived isotopes such as 124I, 89Zr, etc.
Simultaneous imaging of large volume:
  Dynamic imaging of multi-organ systems, multiple lesions
  Kinetic modeling for biologic parameter calculation, e.g., flow, tracer flux

## Slide 7
Spatial resolution in long AFOV scanners
Schmall et al PMB 2016
Transverse | 23 cm | 70 cm
center - 1 cm | 3.96 mm | 4.06 mm
radial 10 cm | 4.63 | 4.77
radial 20 cm | 6.10 | 6.24
Axial | 23 cm | 70 cm
center - 1cm | 4.01 | 4.49
10 cm | 4.18 | 4.77
20 cm | 4.22 | 4.73
~ 0.5-mm loss
Karp, et. al., 2017 IEEE NSS-MIC (Atlanta)

## Slide 8
Lesion detectability with long AFOV scanners
Act.
(μCi/cc)
18cm
600ps
72cm
600ps
72cm
300ps
0.005
0.010
0.025
0.050
0.87 ± 0.04
0.100
0.06 ± 0.04
0.07 ± 0.04
0.60 ± 0.04
0.19 ± 0.04
0.46 ± 0.07
0.79 ± 0.09
0.96 ± 0.05
0.97 ± 0.05
0.73 ± 0.06
0.92 ± 0.03
0.99 ± 0.02
0.99 ± 0.03
Non-DOI, 4x4x20 mm3 crystals, 1 cm lesions with 3:1 uptake, 10 mins. scan time for a 100 cm long phantom
Surti, et. al., PMB 2015,
5343-5358

## Slide 9
Long axial FOV PET: Design Considerations
Performance requirements
  High sensitivity
  DOI measurement no necessary
  Retain, or improve upon state-of-the-art commercial PET/CT – spatial and timing resolution
Scalable platform
  Allow for different axial field-of-view
3 rings 70 cm
1 2 3
6 rings 140 cm
1 2 3 4 5 6
6 rings (w/ spacers) 170-190 cm
1 2 3 4 5 6

## Slide 10
Technology platform: SiPM detectors
3 rings completed May 2018
Philips digital SiPM detector tile: 1-to-1 coupling
  3.86 x 3.86 x 19 mm3 LYSO
Detector module: 28 tiles/module
Detector tile (64 ch)
Detector module (28 tiles)
Detector ring: 18 modules/ring
  23-cm axial length

## Slide 11
Data acquisition and Calibration
Data acquisition for multiple rings
  Common clock (master-slave…)
  Singles events
    Data rate: 100 Mcps/ring
    Data storage: 2 TB/ring
  Time alignment -> off-line sorting
Calibrations Timing and Energy
    Per ring calibration
    Lower temp to reduce dark noise (and deadtime) for improved timing with trigger level 1
< 250 ps @5˚
Notes: -

## Slide 12
Performance Measurements: Count-rate
Chart: Activity Concentration (kBq/cc)
| 44.709742785719698 | 42.51175956231836 | 40.42183176373824 | 37.953307308541142 | 35.860793910701169 | 33.883648912367043 | 31.81440532983034 | 29.683889778856141 | 27.696048480815019 | 23.363001130436601 | 21.122360415192801 | 19.09660950741759 | 16.941824456183529 | 14.841944889428079 | 12.75885032925037 | 10.62793989057962 | 8.4709122280917661 | 6.339352638247516 | 4.235456114045884 | 3.0520742673290449 | 2.1177280570229411 | 1.4974598698299331 | 1.1277336984946451
True (kcps) | 5766.4916 | 5529.2504 | 5297.4317 | 5035.1001 | 4803.7469 | 4573.1978 | 4349.7628 | 4105.851 | 3862.9732 | 3332.1476 | 3053.9688 | 2794.3843 | 2491.6809 | 2219.1832 | 1923.6061 | 1623.2281 | 1305.0448 | 985.8445 | 662.4215 | 479.0975 | 334.8598 | 232.6642 | 176.3206
Scatter (kcps) | 2914.3231 | 2747.8227 | 2612.3703 | 2460.9549 | 2333.1098 | 2210.8549 | 2104.0827 | 1983.0252 | 1861.3069 | 1598.8155 | 1465.2741 | 1334.6781 | 1191.4044 | 1056.1222 | 914.3245 | 769.8547 | 617.2483 | 465.3353 | 311.828 | 225.3851 | 157.2409 | 108.6208 | 82.472
Randoms (kcps) | 18398.7664 | 16646.0339 | 15086.6149 | 13404.1401 | 12038.0352 | 10763.9068 | 9635.3854 | 8476.4194 | 7413.8085 | 5374.4011 | 4469.2545 | 3697.0179 | 2914.3203 | 2290.9943 | 1711.9664 | 1218.4454 | 791.4141 | 460.0112 | 217.7679 | 121.5173 | 65.6603 | 36.5266 | 24.0888
NECR (kcps) | 1227.952 | 1226.6773 | 1220.3111 | 1213.0142 | 1203.448 | 1191.8273 | 1175.969 | 1157.4096 | 1135.8244 | 1077.4202 | 1037.629 | 997.7643 | 941.0477 | 884.7483 | 813.2625 | 729.5719 | 627.6071 | 508.5255 | 368.1173 | 277.8868 | 201.0378 | 143.2794 | 109.901
Inject 15 mCi FDG  5.5 kBq/cc for 70 kg patient
NEMA 70-cm long
Count-rate Phantom
20-cm diameter cylinder with off-center line source
Chart: Activity Concentration (kBq/cc)
| 44.709742785719698 | 42.51175956231836 | 40.42183176373824 | 37.953307308541142 | 35.860793910701169 | 33.883648912367043 | 31.81440532983034 | 29.683889778856141 | 27.696048480815019 | 23.363001130436601 | 21.122360415192801 | 19.09660950741759 | 16.941824456183529 | 14.841944889428079 | 12.75885032925037 | 10.62793989057962 | 8.4709122280917661 | 6.339352638247516 | 4.235456114045884 | 3.0520742673290449 | 2.1177280570229411 | 1.4974598698299331 | 1.1277336984946451
True (kcps) | 5766.4916 | 5529.2504 | 5297.4317 | 5035.1001 | 4803.7469 | 4573.1978 | 4349.7628 | 4105.851 | 3862.9732 | 3332.1476 | 3053.9688 | 2794.3843 | 2491.6809 | 2219.1832 | 1923.6061 | 1623.2281 | 1305.0448 | 985.8445 | 662.4215 | 479.0975 | 334.8598 | 232.6642 | 176.3206
Scatter (kcps) | 2914.3231 | 2747.8227 | 2612.3703 | 2460.9549 | 2333.1098 | 2210.8549 | 2104.0827 | 1983.0252 | 1861.3069 | 1598.8155 | 1465.2741 | 1334.6781 | 1191.4044 | 1056.1222 | 914.3245 | 769.8547 | 617.2483 | 465.3353 | 311.828 | 225.3851 | 157.2409 | 108.6208 | 82.472
Randoms (kcps) | 18398.7664 | 16646.0339 | 15086.6149 | 13404.1401 | 12038.0352 | 10763.9068 | 9635.3854 | 8476.4194 | 7413.8085 | 5374.4011 | 4469.2545 | 3697.0179 | 2914.3203 | 2290.9943 | 1711.9664 | 1218.4454 | 791.4141 | 460.0112 | 217.7679 | 121.5173 | 65.6603 | 36.5266 | 24.0888
NECR (kcps) | 1227.952 | 1226.6773 | 1220.3111 | 1213.0142 | 1203.448 | 1191.8273 | 1175.969 | 1157.4096 | 1135.8244 | 1077.4202 | 1037.629 | 997.7643 | 941.0477 | 884.7483 | 813.2625 | 729.5719 | 627.6071 | 508.5255 | 368.1173 | 277.8868 | 201.0378 | 143.2794 | 109.901
Sensitivity = 45 kcps/MBq
Trues linear over wide range
Scatter fraction 32%
NECR peak ~ 1200 kcps
FDG
@ inject

## Slide 13
CTN Torso Phantom (+ cylinder)
Quantitative Accuracy and detectability for 60 sec scans
360 sec
60 sec
15 sec
30 sec
120 sec
13 mm
10 mm
Lesion detectability (ALROC) measured for embedded lesions using a generalized scan statistic model
Lesion CRC:
(T/B-1)/(Ratio-1)
Lesion precision:
%SD of replicates
Sag
MIP
Tr
Cor
Surti et al, et. al., 2018
IEEE NSS-MIC (Sydney)
Daube-Witherspoon, et. al., 2018
IEEE NSS-MIC (Sydney)

## Slide 14
Studies performed with IRB protocol with informed consent
  CT from commercial PET/CT
    Light restraint of arms and head to aid alignment
    Register (rigid-body) to non-AC image
PennPET Explorer: Human studies
3 Rings: 70 cm axial length
Spatial resolution: 4 mm
TOF resolution: 250 ps
Sensitivity: 45 kcps/MBq
Notes: Last summer we initiated human studies – the true test of a new system.
To date we have performed 10 studies, both static and dynamic, including a couple of clinical patients and research subjects with new tracers. You will see more of these from the next speakers.

## Slide 15
P001: August 15, 2018
F 62 y.o.
163 cm
15 mCi FDG
1 ½ hr p.i.
20 min scan
Human studies
Singles = 15 Mcps/ring
Prompts = 4.7 Mcps
Singles/Prompts = 10:1
Prompts – Delays = 2 Mcps
RF = 60%

## Slide 16
Compare to clinical PET
2 min scan
(sub-sampled)
16 min scan 8 beds
¾ hr p.i.
2 min scan
1 bed
1½ hr p.i.
Clinical TOF PET/CT
PennPET Explorer
Comparable, or better, image quality in shorter scan: scales with sensitivity (7x)

## Slide 17
Reduce scan time
10 min
1¼ min
20 min
10 min
5 min
2½ min
1¼ min
37 sec
Good image quality (and liver uniformity) for 1 min scans

## Slide 18
Delayed scans
Pr-De = 1.0 Mcps
R/T 40%
Pr-De = 85 kcps
R/T = 18%
1½ hr p.i.
3 hr p.i.
20 min
20 min
High sensitivity enables delayed scans with higher contrast

## Slide 19
Delayed scans
7 hr p.i.
19 hr p.i.
Changing Brain-Body kinetics are evident at 10 t1/2’s
30 min
60 min

## Slide 20
Delayed scans
1½ hr
5 hr
9 hr
19 hr
Slow washout of FDG in the myocardium
Post-injection time (minutes)
Myocardium
Brain (Occ)
Brain (L Temp)
0 200 400 600 800 1000 1200
7 hr
9 hr
5 hr
19 hr
3 hr
1½ hr

## Slide 21
Human studies
P003: October 01, 2018
M 79 y.o.
173 cm
15 mCi FDG
Pr-De = 2.6 Mcps
RF = 70%
10 min p.i
3 min
1¾ hr p.i.
20 min
Pr-De = 1.4 Mcps
RF = 50%
4 hr p.i.
30 min
Pr-De = 0.45 Mcps
RF = 30%
Delayed scans: Increased contrast with good statistics
e.g. heart, aorta, rib
Early scans: High data throughput, up to 25 Mcps/ring

## Slide 22
Early scans
Delayed scans
10 min p.i.
3 min
40 min p.i.
3 min
1 ¾ hr p.i.
20 min
4 hr p.i.
30 min
18 hr p.i.
60 min
Clinical scan
70 min p.i.
8 x 1½ min

## Slide 23
PennPET Explorer: Summary to date
Demonstration of design and performance
3-ring, 70-cm AFOV system operational
  Scalable design – can be expanded beyond 3 rings
  Low temperature at tiles without condensation
  High data throughput – parallel singles acquisition
  Data correction and image reconstruction implemented
Intrinsic performance measured
  4-mm spatial resolution
  250 ps timing resolution
  10% energy resolution
Human Studies
  Excellent image quality
  Demonstration of short scans and delayed imaging
    Interesting kinetics, even for FDG
  Dynamic imaging performed – inject on table
Notes: Studies to optimize the trade-offs in scanner design – different for large FOV compared to clinical systems
TOF benefit most appreciated for low activity (or low dose)

## Slide 24
Long Axial FOV PET: Steps to optimize for clinical research
  PennPET Explorer
    Expand from 3 rings to 6 rings
    Integrate with commercial CT in new imaging facility
    Demonstration of benefits of high sensitivity
      Shorter scans, lower dose – optimize scan protocols
      Delayed scans (slow biology), dynamic scans (fast biology)
    Validation of long axial FOV design
      Data acquisition – count-rate requirements, optimize acceptance angle
      Data correction/Image reconstruction methods – quantitative accuracy
    Test and evaluate new applications
      Clinical and research
Notes: There are several immediate tasks to complete the PennPET Explorer – it is currently being upgraded from 3 to 6 rings, without gaps, and will be integrated with a CT in a new facility in the next 2 months.
The validation of the design will continue once it is operational again, in an effort to optimize the performance of our instrument, and also address more general questions for long axial FOV PET.
This includes demonstration of the benefits of high sensitivity and larger axial coverage, as applied to various tracers that are used to study both fast and slow biology. More will be said about this in the following talks

## Slide 26
Data correction & Image reconstruction
Uniformity test: meter-long cylinder
Data correction
  Efficiency normalization for all lines-of-response
  Rotating line source at 30-cm radius
  Easy to extend to longer axial FOV
Rotating line source in carbon fiber tube
Image reconstruction
  List-mode TOF-OSEM
  Extended to handle larger data/image structures
  Overlapping blob basis functions constrain noise while preserving signal; basis functions are then interpolated into image voxels after reconstruction
  PennPET: blob (B) into 2-mm voxels (for body)
z
Notes: -

## Slide 27
Currently: 5 of 7 rows read out
16.4 cm axial length: 6.5 cm gap
Current data includes gaps in axial direction
Uniform image and noise profile – within 90% axial FOV
Firmware update will increase ring length to 22.9 cm
  3 rings  70 cm, 6 rings  140 cm
  Potential to separate rings for longer axial FOV
Axial Gaps between rings
Reconstructed pipe phantom image
Axial noise profile in pipe phantom image
Ring 1
2
3
Ring 1
2
3
64 cm

## Slide 28
Reconstruction: Body: 2mm3 voxels
Brain: 1-mm3 voxels
Optimize reconstruction

## Slide 29
Human studies
P002: September 26, 2018
F 56 y.o.
155 cm
15 mCi FDG
5 hr p.i.
15 min
Washout of FDG varying with organ is evident in delayed scan (as is lunch)
1½ hr p.i.
10 min

## Slide 30
Maximum axial acceptance angle (center FOV) ±42˚
Maximum axial acceptance angle (for brain) ±8˚
Impact of axial location
Larger axial acceptance Improved statistics
Resolution degradation not evident