## Slide 1: Positron emission tomography imaging using polarization-correlated annihilation quanta – an experimental study
Ana Marija Kožuljević1, Tomislav Bokulić1, Darko Grošev2, Zdenka Kuncic3, Siddharth Parashari1, Luka Pavelić4, Marijan Žuvić2 and Mihael Makek1
1 Department of Physics, Faculty of Science, University of Zagreb, Zagreb, Croatia
2 Department of Nuclear Medicine and Radiation Protection, University Hospital Centre Zagreb, Zagreb, Croatia
3 School of Physics, The University of Sydney, Sydney, NSW, Australia
4 Institute for Medical Research and Occupational Health, Zagreb, Croatia
5th Jagiellonian Symposium on Advances in Particle Physics and Medicine
June 29 – July 7, 2024

## Slide 2: Motivation
Two gammas from positronium annihilation have:
      511 keV energy
      opposite momenta
      orthogonal polarizations

## Slide 3: Motivation
Two gammas from positronium annihilation have:
      511 keV energy
      opposite momenta
      orthogonal polarizations
Implemented in PET
YES
YES
NO

## Slide 4: Motivation
Two gammas from positronium annihilation have:
      511 keV energy
      opposite momenta
      orthogonal polarizations
Implemented in PET
YES
YES
NO
Polarization correlations could be used as an additional handle to reduce background in PET
Improved image quality
*McNamara A et al. (2014) Phys. Med. Biol., 59: 7587; Toghyani M et al. (2016) Phys. Med. Biol., 61: 5803; Watts DP et al. (2021) Nat. Commun., 12: 2646;
Kim D et al. (2023) JINST, 18(07):P07007

## Slide 5: Azimuthal correlations of annihilation quanta
Double Klein-Nishina differential cross section:​

## Slide 6: Azimuthal correlations of annihilation quanta
Double Klein-Nishina differential cross section:​
Maximum at |φ1-φ2|=90°

## Slide 7: Azimuthal correlations of annihilation quanta
Double Klein-Nishina differential cross section:​
Maximum at |φ1-φ2|=90°
Polarimetric modulation factor:

## Slide 8: Azimuthal correlations of annihilation quanta
Double Klein-Nishina differential cross section:​
Maximum at |φ1-φ2|=90°
Polarimetric modulation factor:
Δφ [°]
Max for Θ1= Θ2= ~82°:
μ = 0.48
Background events will lack such correlation!
Double K-N function with polarimetric modulation factor 0.48
a.u.
Δφ [°]

## Slide 9: Laboratory setup
Video made by Ena Piškorić
Pixelated Compton polarimeters in one layer

## Slide 10: Laboratory setup
Scintillating crystals
  GaGG:Ce
    8.1% ± 0.5% at 511 keV
Silicon Photomultiplier (SiPM)
(Hamamatsu Photonics, Japan, model S13361-0808AE), 1:1 coupling
ToFPET2 ASIC read-out system
Parashari et al., Nucl. Instrum. Methods Phys. Res. A (2022) 167186
Pixelated Compton polarimeters in one layer

## Slide 11: Data selection
EVENT
1 pixel fired per module with 511±3σ keV deposited energy
Selection of events with photoelectric interaction
2 pixels fired per module

## Slide 12: Data selection
1 pixel fired per module with 511±3σ keV deposited energy
Selection of events with photoelectric interaction
2 pixels fired per module
elliptical cut with sum of deposited energies 511±3σ keV
Counts
EVENT
b
First pixel -
lower energy deposition!
True for ~55% of the events
(A. M. Kožuljević et al., Condensed Matter (2021), 6, 43)
a
Compton scattering events
Selection of events with two pixels fired with sum of deposited energies 511±3σ keV
Selection of Compton events with two pixels fired with sum of deposited energies 511±3σ keV

## Slide 13: Data selection
EVENT
2 pixels fired per module
1 pixel fired per module with 511±3σ keV deposited energy
elliptical cut with sum of deposited energies 511±3σ keV
Θ angle reconstruction
Selection of events with photoelectric interaction
Counts
Compton scattering events
Distribution of the selected events' Compton scattering angles Θ
θ [° ]

## Slide 14: Data selection
EVENT
2 pixels fired per module
1 pixel fired per module with 511±3σ keV deposited energy
elliptical cut with sum of deposited energies 511±3σ keV
Θ angle reconstruction
Selection of events with photoelectric interaction
Counts
θ [° ]
ϕ angle reconstruction
Compton scattering events
Distribution of the selected events' Compton scattering angles Θ

## Slide 15: Data selection
EVENT
2 pixels fired per module
1 pixel fired per module with 511±3σ keV deposited energy
elliptical cut with sum of deposited energies 511±3σ keV
Θ angle reconstruction
Selection of events with photoelectric interaction
Counts
θ [° ]
ϕ angle reconstruction
Compton scattering events
Selection of events with correlations in polarizations
Azimuthal angle difference: ϕ 1- ϕ 2

## Slide 16: Polarization correlations - reconstruction
After the acceptance correction
Distribution of the azimuthal angle differences
after both gammas undergo Compton scattering
(from measurements)
Blue – measured data
Red – Klein-Nishina fit function
GAGG:Ce 2.9 mm
A. M. Kožuljević et al., Condensed Matter (2021), 6, 43
An example of N(ϕ1- ϕ2 ) distribution for correlated gamma photons in GAGG detector module for Θ= 82°, compared to the simulated distribution for non-correlated photons.
Acceptance correction:

## Slide 17: Polarization correlations - addendum
Experimental setup for measuring polarization correlations after one of the annihilation photons undergoes previous scatter
Parashari et al. (2024) Physics Letters B, 852:138628

## Slide 18: Polarization correlations - addendum
Azimuthal angle difference distributions observed for different Compton scattering angles of the scattered annihilation photons
Experimental setup for measuring polarization correlations after one of the annihilation photons undergoes previous scatter
Polarization correlations observed even after the scatter of one of the annihilation gammas
Parashari et al. (2024) Physics Letters B, 852:138628

## Slide 19: Polarization correlations - addendum
Azimuthal angle difference distributions observed for different Compton scattering angles of the scattered annihilation photons
Experimental setup for measuring polarization correlations after one of the annihilation photons undergoes previous scatter
Polarization correlations observed even after the scatter of one of the annihilation gammas
Limits the ability of the method to reduce background by eliminating scatter coincidences
Parashari et al. (2024) Physics Letters B, 852:138628
The method is still applicable for random and multiple coincidences

## Slide 20: The PET Demonstrator
Four super-modules of segmented scintillators
  16 x 16 pixels in each module
Diameter range: 420 – 700 mm​
Precise rotation around the scanner axis​

## Slide 21: Preliminary results
Two Ge-68 line sources in aluminum encapsulation
Epoxy phantom, 3 cm in diameter
Data acquired with 3.2 mm matrix pitch GaGG:Ce modules
Diameter: | 430 mm
Number of positions (angle): | 12 (15 deg)
Time of acquisition (per position): | ~ 2.3 h
Activity (each source): | 45.5 MBq
Distance between the sources: | ~ 2 cm

## Slide 22
Images reconstructed from events which are selected
as the ones with polarization correlations where
Θ∈[72,90], |Δϕ|∈[70, 110]
Images reconstructed from events which are selected as the single-pixel (PE) events
Extended sources (2 Ge-68 sources in stainless steel, 45.5 MBq each)
a & d - coronal view
b & e – sagittal view
c & f - axial view
OMEGA software (MATLAB), OSEM - NO CORRECTIONS!
128x128 pixels; FOV: 20x20 cm2, 31 slices; OSEM: 10 iterations, 6 subsets
RAW IMAGES!
Preliminary results

## Slide 23: Study of the spatial resolution of the novel PET device
Image profile from single-pixel (PE) events
Image profile from events with polarization correlations where Θ∈[72,90], |Δϕ|∈[70, 110]
Preliminary results

## Slide 24: Study of the spatial resolution of the novel PET device
Image profile from single-pixel (PE) events
Image profile from events with polarization correlations where Θ∈[72,90], |Δϕ|∈[70, 110]
Preliminary results
LOR ambiguity

## Slide 25
NEMA NU-4 Small Animal Phantom
a)
b)
c)
d)
a) and b): Photoelectric events
128x128 pixels
FOV: 200
c) and d): Polarization correlated events
75x75 pixels
FOV: 195
Initial activity: ~ 400 MBq Ga-68
Number of positions: 12 (15 deg)
Total acquisition length: ~ 1h
3.2 mm detector matrix pitch
Preliminary results
b) and d)
a) and c)

## Slide 26
u.
u.
Signal to random background ratio (from coincidence time spectra)
Counts
Counts
Coincidence time spectrum from events with polarization correlations where Θ∈[72,90], |Δϕ|∈[70, 110]
Coincidence time spectrum from single-pixel (PE) events
Counts
Counts
Gaussian fit to the peak

## Slide 27
SBR = (Atot-ABG)/ABG​
Total signal (|Δt| < mean ± 3σ)
Background signal
(Left: Δt < mean - 15σ and Δt > mean - 21σ,
Right: Δt > mean + 15σ and Δt < mean + 21σ)
Counts
Counts
Signal to random background ratio (from coincidence time spectra)
Coincidence time spectrum from events with polarization correlations where Θ∈[72,90], |Δϕ|∈[70, 110]
Coincidence time spectrum from single-pixel (PE) events
Counts
Counts

## Slide 28: Signal to random background ratio (from coincidence time spectra)
Preliminary results for 2 line Ge-68 sources in a phantom, 430 mm diameter, total activity ~90 MBq
u.
u.
Selection of events with polarization correlations where Θ∈[72,90], |Δϕ|∈[-20, 20]
Selection of events with polarization correlations where Θ∈[72,90], |Δϕ|∈[70, 110]
SBR
Position #
SBR vs. Data-acquisition position on the ring for 2 line Ge-68 sources with phantom at 430 mm diameter. Comparison between single-pixel events, and events with polarization correlations at selected |Δϕ|.

## Slide 29: Signal to random background ratio (from coincidence time spectra)
Preliminary results for 2 line Ge-68 sources in a phantom, 430 mm diameter, total activity ~90 MBq
u.
u.
Position #
Position #
SBR
SBR
SBR vs. Data-acquisition position on the ring for 2 line Ge-68 sources with phantom at 430 mm diameter. Comparison between single-pixel events, and events with polarization correlations at selected pixel distances d.
SBR vs. Data-acquisition position on the ring for 2 line Ge-68 sources with phantom at 430 mm diameter. Comparison between single-pixel events, and events with polarization correlations at selected Θ.

## Slide 30: Signal to random background ratio (from coincidence time spectra)
Preliminary results for NEMA phantom filled with Ga-68, 430 mm diameter
u.
u.
SBR
Position #
SBR vs. Data-acquisition position on the ring for NEMA phantom at 430 mm diameter. Comparison between single-pixel events, and events with polarization correlations at selected |Δϕ|.
Initial activity ~400 MBq, final activity ~200 MBq
SBR vs. Data-acquisition position on the ring for NEMA phantom at 430 mm diameter. Comparison between single-pixel events, and events with polarization correlations at selected Θ.
SBR
Position #
SBR
Position #

## Slide 31: Conclusions
It is possible to measure polarization correlations of the annihilation quanta using single-layer Compton polarimeters.
We have demonstrated that image reconstruction solely from polarization correlated annihilation quanta is possible.
  Limiting factor: LOR ambiguity
Signal to random background ratio from coincidence time spectra of correlated events selected for Θ∈[72,90], |Δϕ|∈[70, 110] exhibits higher values when compared to events with PE and other choices of Θ and |Δϕ|.

## Slide 32: Thank you!
This work was supported by the “Research Cooperability” Program of the Croatian Science Foundation funded by the European Union from the European Social Fund under the Operational Program Efficient Human Resources 2014–2020,
Grant number PZS-2019-02-5829.

## Slide 33: Signal to random background ratio (from coincidence time spectra)
Preliminary results for 2 line Ge-68 sources in a phantom, 620 mm diameter
u.
u.
Position #
Position #
SBR
SBR
SBR vs. Data-acquisition position on the ring for 2 line Ge-68 sources with phantom at 620 mm diameter. Comparison between single-pixel events, and events with polarization correlations at selected |Δϕ|.
SBR vs. Data-acquisition position on the ring for 2 line Ge-68 sources (no phantom) at 620 mm diameter. Comparison between single-pixel events, and events with polarization correlations at selected |Δϕ|.