## Slide 1
Siddharth Parashari*,‡, Tomislav Bokulić*, Damir Bosnar*, Ana Marija Kožuljević*, Zdenka Kuncic†, Luka Pavelić#, Petar Žugec* and Mihael Makek*
*Department of Physics, Faculty of Science, University of Zagreb
†School of Physics, University of Sydney
#Institute for Medical Research and Occupational Health
Development of Polarization-Sensitive Positron Emission Tomography Demonstrator based on Single-layer gamma-ray polarimeters
‡siddharth@phy.hr
13-Jul-2022, Krakow
4th Jagiellonian Symposium on Advances in Particle Physics and Medicine

## Slide 2
Photon polarization in positron annihilation event
Positron annihilation may result in two entangled and orthogonally polarized gamma photons.
It is found that the polarization correlations can be utilized as an additional handle to improve Signal to Noise Ratio (SNR) which may improve medical imaging with Positron Emission Tomography (PET).
Introduction
How to measure gamma-ray polarization in PET ?
The gamma polarization is related to the azimuthal angle in the Compton scattering process, so the initial correlation of polarization translates to the correlation of azimuthal angles in true coincidence events, which is not present in the background.
Compton scattering
Question ?
How to implement polarization measurement in PET cost efficiently?
4th Jagiellonian Symposium on Advances in Particle Physics and Medicine
2
[Z. Kuncic et al., NIM A648 (2011), A.L. McNamara, et. al., Phys. Med. Bio. 59 (2014) 7587, M. Toghyani et. al., Phys. Med. Bio. 61 (2016) 5803, D. Watts et. al., Nat. Commun. 12 (2021) 2646]

## Slide 3
Klein-Nishina differential cross-section for scattering of linearly polarized gamma photon
Arbitrary Unit
Gamma polarization measurement via Compton scattering
where θ is the scattering angle and φ is the angle between the scattering plane and the polarization vector
Gamma is most likely to be scattered at azimuthal angle φ perpendicular to the polarization vector (cos φ = 0)
The sensitivity to polarization is the largest for scattering at θ = 90°
Polarization is correlated to the azimuthal scattering angle φ.
Polarization correlations in paired Compton events
The cross-section for scattering of two linearly polarized γ-particles is given by
The cross-section has maxima when ; keeping θ fixed.
4th Jagiellonian Symposium on Advances in Particle Physics and Medicine
3
The Polarimetric Modulation Factor μ is defined as,
μ reaches maximum μ = 0.48 when θ1 = θ2 = 82o
Φ (deg)
Arbitrary Unit
Arbitrary Unit
ΔΦ (deg)
[Pryce and Ward, Nature 160 435; Snyder, Pasternak, Hornbostel, Phys. Rev. 73 (1948) 440-8]
[M. Toghyani et. al., Phys. Med. Biol. 61 (2016) 5803]

## Slide 4
Motivation
Modern PET detectors : Highly segmented, large coverage
Most common setups
to detect azimuthal :
polarization
Detectors with 2 senisitive layers – 1st for measuring the recoil electron, 2nd for the scattered gamma
Detection of recoil electron and scattered photon
A PET system based on 2 (or more)-layer detectors would dramaticaly increase the cost of the apparatus
Drawbacks
Feasibility of the measurement of polarization correlations using single layer detectors
Single layer detectors
M. Makek et. al., Nucl. Instr. Meth. A 958 (2020) 162835.
Proof of concept using 4x4 LFS crystals 3 x 3 x 20 mm3
Energy resolution of detectors determines θ precision,
Segmentation and material determine φ resolution and acceptance
1. Improving energy resolution Choosing suitable detector material
2. Improving azimuthal resolution Small pixel dimensions in a detector
4th Jagiellonian Symposium on Advances in Particle Physics and Medicine
4
μ = 0.29 ± 0.04; 72o<θ1,2 <90o

## Slide 5
Comparison of different setups for measurements of gamma polarization correlations
Fig. 2. Individual pixel resolution (%) at 511 keV in the GaGG 1.9 mm detector modules.
Pixel Array – 13
Pixel Array – 14
4th Jagiellonian Symposium on Advances in Particle Physics and Medicine
5
Fig. 1. Schematic diagram of experimental setup – example for GaGG:Ce 8x8 matrix
TOFPET2 DAQ system from PETSys
S. Parashari et al.,“Optimization of detector modules for measuring gamma-ray polarization in PET”, [Accepted in NIM A.] PET”, arXiv:2205.0107]

## Slide 6
Data analysis
The Compton events in each module are selected requiring that exactly two pixels fire [Fig. 3(a)], that the energy deposited in the module is within 511 keV±3σ [Fig. 3(b)] and that pixel energies correspond to Compton kinematics, [Fig. (c), (d)].
Selection of Compton Events –
For ex. Pixels 10 & 12 fired in a Compton Event
Sum of Pixel Energies
Two Pixel fired per event
Energy deposition in Selected Pixel
Selected Compton Events
(Color Region)
( a )
( b )
( c )
( d )
Fig. 3. (a) Number of pixel fired in all events; (b) sum of pixel energies in Compton events; (c) energy deposited in an individual pixel fired in Compton event; (d) selected Compton events
4th Jagiellonian Symposium on Advances in Particle Physics and Medicine
6
10
12

## Slide 7
[A. M. Kožuljević et. al., Condensed Matter 2021, 6, 43]
Data analysis
Reconstruction of Compton Scattering Angles –
4th Jagiellonian Symposium on Advances in Particle Physics and Medicine
7
Detector Acceptance –
The acceptance-corrected φ distribution is obtained as
where An is obtained with f1 and f2 from different (uncorrelated) events.
+
The Scattering angle reconstruction,
Reconstruction of azimuthal angles,
(a)
GaGG 3(mm) x 3(mm)
Fig. 4. Reconstructed (a) 𝜃 and (b) 𝝓1- 𝝓2 (correlated vs uncorrelated)
(b)

## Slide 8
Polarization correlations of gamma-ray from positron annihilation
Distributions : For various inter-pixel distances d1,2 > 4.5 mm for different angular range in 72o < 𝜃1,2 < 90o .
Modulation factor μ extracted by fit :
Fig 5. Observed azimuthal difference distributions for (a) GaGG_3.0mm (b) GaGG_2.9mm (c) LYSO_2.0mm (d) LYSo_1.9mm &
(e) GAGG_1.9mm for 72o < 𝜃1,2 < 90o
(a)
(b)
(c)
(e)
(d)
GaGG 3.0 mm
GaGG 2.9 mm
LYSO 2.0 mm
GaGG 1.9 mm
LYSO 1.9 mm
8
[S. Parashari et al., PET”, arXiv:2205.0107]

## Slide 9
Modulation factors in LYSO 1.9 mm
at different scattering angles
The dependence of the modulation factor, 𝜇, on the azimuthal resolution <Δφ> at different 𝜃1,2 is explored in each detector configuration. An examples of obtained Modulation factors, 𝜇 with azimuthal resolution <Δφ> for LYSO 1.9 mm detector is shown in figure.
<Δφ>o
Fig. 6. Modulation (%) vs azimuthal resolution <Δφ> at different angular ranges for LYSO 1.9 mm detector.
4th Jagiellonian Symposium on Advances in Particle Physics and Medicine
9
Observations –
Larger modulation factors are observed for angular ranges closer to 82o in comparison to 70o.
Rising modulation amplitude is observed with lower Azimuthal Resolution <Δφ> or larger inter-pixel distances d(mm).
We have achieved a well pronounced polarimetric performance of finer segmented detector modules.
[S. Parashari et al.,“Optimization of detector modules for measuring gamma-ray polarization in PET”, [Accepted in NIM A.] PET”, arXiv:2205.0107]

## Slide 10
Comparison of μ vs Δφ among modules at different scattering angles
We compared the polarimetric performance μ of individual detectors with azimuthal resolution <Δφ> at different angular ranges.
<Δφ>o
<Δφ>o
Fig. 7. Comparison of modulation (%) vs azimuthal resolution <Δφ> for all detectors at angular ranges (a) 72o < 𝜃1,2 < 92o and (b) 77o < 𝜃1,2 < 87o
(a)
(b)
4th Jagiellonian Symposium on Advances in Particle Physics and Medicine
10
Observations –
Modulation factors increases further for
77o < 𝜃1,2 < 87o as we select events closer to
scattering angle 82o.
[S. Parashari et al.,“Optimization of detector modules for measuring gamma-ray polarization in PET”, [Accepted in NIM A.] PET”, arXiv:2205.0107]

## Slide 11
Fig. 8. Modulation amplitude vs. Δ𝜃 at scattering angles 77o<𝜃1,2<87o and ⟨Δ𝜙⟩ ≈ 16 ± 1o for all configurations. The systematic uncertainty is shown by the blue boxes.
Comparison of μ vs Δθ among modules at scattering angle 77o<θ1,2<87o
We compared the polarimetric performance μ of individual detectors with scattering angle resolution Δθ at scattering angle 77o<𝜃1,2<87o and ⟨Δ𝜙⟩ ≈ 16 ± 1o.
4th Jagiellonian Symposium on Advances in Particle Physics and Medicine
11
Observations –
Energy Resolution of the detectors translates into the uncertainty in the scattering angle θ, which results in the under-estimation of the modulation factors.
[S. Parashari et al.,“Optimization of detector modules for measuring gamma-ray polarization in PET”, [Accepted in NIM A.] PET”, arXiv:2205.0107]

## Slide 12
Development of PET Demonstrator
4th Jagiellonian Symposium on Advances in Particle Physics and Medicine
12
Four super-modules16 x 16 scintillator pixels each
Diameter range: 420 –700 mm
Rotation around the scanner axis
Super-Modules
(16x16 pixels)

## Slide 13
Summary
4th Jagiellonian Symposium on Advances in Particle Physics and Medicine
13
Angular correlation of annihilation quanta were successfully measured with the single-layer pixelated scintillation detectors (cost-efficient scalability to larger systems).
Polarimetric performance of 5 module types from 1.9 – 3.0 mm pixel sizes were assessed.
Higher modulation was obtained with finely segmented pixel detectors.
The measured modulation amplitude shows a rising trend with better azimuthal resolution.
The dependance on the energy resolution of the material is not pronounced, LYSO and GAGG perform comparably.
Next: Commission of the demonstrator and test its performance with extended sources.

## Slide 14
This work is supported by the “Research Cooperability” Program of the Croatian Science Foundation co-funded by the European Union from the European Social Fund under the Operational Programme Efficient Human Resources 2014–2020, Grant number PZS-2019-02-5829.
Thank you for your attention !!!
4th Jagiellonian Symposium on Advances in Particle Physics and Medicine
14

## Slide 16
Comparison of modulation factors in detector modules
at different scattering angles (θ ± dθ)
The dependence of the modulation factor, 𝜇, at different 𝜃1,2 ranges is explored in each detector configuration.
The obtained modulation factors, 𝜇 for 72o < 𝜃1,2 < 90o and 77o < 𝜃1,2 < 87o and 80o < 𝜃1,2 < 84o at inter-pixel distance d > 4.5 (mm) is shown for all modules.
Larger modulation are observed for 80o < 𝜃1,2 < 84o as compared to 72o < 𝜃1,2 < 90o angular range, which is expected from the theory [ Nature, vol. 160, Sep. 1947 ]
S. Parashari et al.,“Optimization of detector modules for measuring gamma-ray polarization in PET”, [Accepted in NIM A.] PET”, arXiv:2205.0107]
Better Modulations are achieved within a narrower angular range 80o < 𝜃1,2 < 84o i.e. closer to the maxima condition .
4th Jagiellonian Symposium on Advances in Particle Physics and Medicine
16

## Slide 17
Sydney, 21/6/2022
The SiLGaP project
17
70° < q1,2 < 90°
70° < q1,2 < 90°
60° < q1,2 < 80°
60° < q1,2 < 80°
GAGG-2.9
LYSO-2.0
Measurement
Fit to data
Simulation
Measurement
Fit to data
Simulation
Measurement
Fit to data
Simulation
Measurement
Fit to data
Simulation
Comparison of the experiment data with simulations
M. Makek et. al., 2021 Virtual IEEE Nuclear Science Symposium And Medical Imaging Conference

## Slide 18
Which of the two pixels fired first?
Ambiguous for θ > 60°
g
511 keV gamma Compton energies
The “θ – ambiguity”
4th Jagiellonian Symposium on Advances in Particle Physics and Medicine
18
Notes: Correct the footer

## Slide 19
The cross section favors forward angles
Backward scattered gammas, θ >90°:
  Lower energy
  Shorther absorption length
Always assume forward scenario
(i.e. the pixel with lower energy corresponds to e-)
4th Jagiellonian Symposium on Advances in Particle Physics and Medicine
19
The “θ – ambiguity”
160 140 120 100 80 60 40 20 0
qSim (deg)
qReco (deg)
SIMULATION
[A. M. Kožuljević et. al., Condensed Matter 2021, 6, 43]

## Slide 20: The „q-ambiguity”
Event topology helps: events with larger inter-pixel distances d have a higher fraction of correctly determined angles
Implication on LOR determination:
  30% correct pixel in both modules
  50% correct pixel in only one module
  20% incorrect LOR
Table: the fraction of correctly determined q in GAGG, assuming forward scattering
[A. M. Kožuljević et. al., Condensed Matter 2021, 6, 43]
LORs
4th Jagiellonian Symposium on Advances in Particle Physics and Medicine
20